Webb Research Corporation

82 Technology Park Drive, E. Falmouth, MA 02536-4441   (508) 548-2077  FAX (508) 540-1686

USER MANUAL – APEX-SBE PROFILER

APEX-SBE INSTRUMENTS

Serial numbers:  52 - 54

REV DATE: 5/28/99

BSH Koltermann, Contract no.  1055/M54/98/10

WRC Job no. 625

I.	ALKALINE BATTERY WARNING	2

II.	Reset and Self Test	2

III.	Mounting Damper Disk	3

IV.	Deployment	3

V.	ARGOS DATA	3

A.	SERVICE ARGOS PARAMETERS	3

B.	DATA FORMAT	4

C.	TEST MESSAGE FORMAT	5

D.	CRC	6

E.	CONSTANTS	7

VI.	MISSIONS	9

VII.	RECORDS & CALIBRATIONS	11

# ALKALINE BATTERY WARNING

The profiler contains alkaline "D" cells.

There is a small but finite possibility that batteries of alkaline cells will release a combustible gas mixture.  This gas release generally is not evident when batteries are exposed to the atmosphere, as the gases are dispersed and diluted to a safe level.  When the batteries are confined in a sealed instrument mechanism, the gases can accumulate and an explosion is possible.

Webb Research Corp. has added a catalyst inside of these instruments to recombine Hydrogen and Oxygen into H2O, and the instrument has been designed to relieve excessive internal pressure buildup by having the upper endcap release.

Webb Research Corp. knows of no way to completely eliminate this hazard.  The user is warned, and must accept and deal with this risk in order to use this instrument safely as so provided.  Personnel with knowledge and training to deal with this risk should seal or operate the instrument.  Webb Research Corp. disclaims liability for any consequences of combustion or explosion.

# Reset and Self Test

Profilers are shipped to the deployment site in Hibernate mode.  Shortly before deployment, reset the profiler by passing a magnet over the marked location on the pressure case.  The profiler will run a self-test, transmit for 6 hours with the bladder extended, and then begin its pre-programmed mission.

The six ARGOS transmissions during self-test and the transmissions during the initial 6 hour period contain data about the instrument and are outlined in  (V) ARGOS DATA, part (C) TEST MESSAGE FORMAT.

Procedure:

- Hold the provided magnet at RESET position marked on the hull for several seconds.

- Note:  The internal magnetic reed switch must be activated (held) for at least one second to reset the instrument.  (This is to provide a safety against accidental reset during transport.)  Thus, if the ALACE does not respond as below, the instrument was probably not reset.

- The air pump will operate for 1 second.

- The PTT will transmit 6 times at 6 second intervals.  Place the ARGOS receiver/beeper close to the antenna to detect transmissions.

- The piston pump will begin to operate.  The piston will move to the retracted Ballast Position, if not already there, pause 2 seconds and then move to full extension.

- The bladder will expand, this should take 15 - 25 minutes.

- After the piston pump stops the air pump will come on and inflate the air portion of the bladder taking 20 - 30 seconds.

- The PTT will transmit at the mission specified ARGOS rate.

- 6 hours after reset, transmissions will cease, the piston pump will retract and the bladder will deflate, the profiler begins its programmed mission.

During self-test, the controller checks the internal vacuum sensor.  If the internal pressure has increased above a preset limit (i.e. hull leakage caused loss of vacuum), the instrument will not pump.

If you do not detect the 6 test transmissions, and if the bladder does not inflate, then the self-test has failed and the instrument should not be deployed.

# Mounting Damper Disk

To aid surface following, a square disk is pre-mounted to the outside of the pressure case using 2 pvc collars, an O-ring, and titanium hardware.

# Deployment

- RESET instrument.

- SELF-TEST starts automatically (see above).

- When piston pump stops, air pump inflates, external bladder is full, PTT will transmit for 6 hours at ARGOS Repetition rate intervals.  Normally 90 seconds.

- Six hours after reset, the piston pump will retract and bladder will deflate.

- Pass a rope through the hole in the stability disk.

- Holding both ends of the rope, carefully lower the ALACE into the water.

- Take care not to damage the antenna.

- Do not leave the rope with the instrument, release one end and retrieve the rope.

- The ALACE will remain on the surface until the 6 hour interval has expired.

# ARGOS DATA

## SERVICE ARGOS PARAMETERS

The user must specify various options to Service ARGOS.  These choices depend on how the user wishes to receive and process data.  Typical parameters are listed below:

- Standard location.

- Processing: Type A2 (pure binary input; hexadecimal output)

- Results Format: DS (all results from each satellite pass), Uncompressed.

- Distribution Strategy:  Scheduled, all results, every 24 hours.

- Number of bits transmitted:  32

## DATA FORMAT

Data is sent via ARGOS in 32 byte hex messages.  The number of 32 byte messages sent depends on the programmed quantity of temperature measurements per profile.  (See section (V) ARGOS DATA, part (E) CONSTANTS.)

Format for message number 1 only:

Byte #

- 01  CRC, described in section C.

- 02  Message number, the total number of messages for each float data set is given in section D.  Messages are transmitted in sequential order starting with 1 and incrementing by one for 	the data set.

- 03  Message block number, begins as 1 and increments by one for every ARGOS message 	data set.  This, combined with the ARGOS repetition rate (section VI), allows the user to 	track surface drift.  Byte 03 will roll-over at 256 and will reset to 1 on each new profile.

- 04 & 05  Serial number, identifies the controller board number.  (This may not be the same 	as instrument number.)

- 06  Profile number, begins with 1 and increases by one for every float ascent.

- 07  Profile length, is the number of six byte STD measurements in the profile.  Total 	number of bytes of STD data from each profile depends on the sampling strategy 	chosen.  This is described in section (V) ARGOS DATA, part (E) CONSTANTS.

- 08  Profile termination flag byte, can have the following values (hex):

00	Pressure reached surface pressure.

02	Pressure reached zero.

- Pressure unchanged for 25 minutes.  (Does not terminate profile.)

- Piston fully extended before surface

- UP time expired before surface and was reset.

- 09  Piston position, recorded as the instrument reaches the surface.

- 10 & 11  Bottom temperature, sampled just before instrument begins ascent.

- 12 & 13  Bottom salinity, sampled just before instrument begins ascent.

- 14 & 15  Bottom pressure, sampled just before instrument begins ascent.

- 16  Battery voltage, nominally at 15 volts and decreases throughout the life of the float.

- 17 & 18  Surface pressure, as recorded just before last descent with an offset of +5db.

- 19  Internal vacuum, as recorded just before last descent.

- 20  Bottom piston position, the linear pot count recorded at the target depth.

- 21 to 32  6 bytes in sequence:

2 bytes temperature

2 bytes salinity

2 bytes pressure

Format for message number 2 thru 12 that follow:

Byte #

- 01  CRC, described in section C.

- 02  Message number

- 03 to 32  6 bytes in sequence:

2 bytes temperature

2 bytes salinity

2 bytes pressure

## TEST MESSAGE FORMAT

The test message is sent whenever an I2 command is given, the six transmissions during the startup cycle, and during the six hour surface mode period prior to the first dive.  Each test message has 32 bytes, in hex unless otherwise noted, with the following format:

Byte #

- 01  CRC, described in section C.

- 02  Message number, always 01.

- 03  Message block number, begins as 1 and increments by one for every ARGOS message.

- 04 & 05  Serial number, identifies the controller board number.  (This may not be the same as 	instrument number.)

- 06  Profile number, always 00.

- 07  Message type flag, 20 for test message, 40 for 6 hour surface message.

- 08 & 09  Current pressure.

- 10  Battery voltage, nominally at 15 volts.

- 11  Internal vacuum, nominally at 5 inches Hg.

- 12  Piston Position, a count value 12-248.

- 13 Float status byte, 08 for float up, 20 piston fully extended, 40 piston running.

- 14  Hour, the following is the time from startup (in decimal).

- 15  Minutes.

- 16  Seconds.

- 17  ARGOS repetition rate constant (multiply by 2, add 6 = nominally 90  6 seconds).

- 18  Up time, intervals.

- 19  & 20  Down time, intervals.

- 21  Trip interval time, hours.

- 22 & 23  Target pressure, in dbars.

- 24  Target piston position, in counts.

- 25  Depth correction factor, in counts.

- 26  Ballast piston position, normally 24 counts.

- 27  Fully extended piston position, nominally 248 counts.

- 28  OK vacuum count, nominally 2 inches Hg.

- 29  Initial piston extension, in counts.

- 30  Month, software version number (in decimal).

- 31  Day, software version number (in decimal).

- 32  Year, software version number (in decimal).

## CRC

Because ARGOS data may contain transmission errors, the first byte of each message contains an error checking value.  This value is a Cyclic Redundancy Check (CRC), and is calculated as a function of the message content (bytes 2 to 32).

- For each message, calculate a CRC value

- Compare the calculated CRC to the transmitted CRC (byte no. 1)

- If the calculated and transmitted CRC values are not equal, the message has been corrupted and should be deleted before further data processing.

Below is a sample program (in BASIC) to calculate the CRC value for a message.  This program can be provided upon request in Basic, Fortran or C.

DECLARE FUNCTION CRC% (IN() AS INTEGER, N AS INTEGER)

‘CRC routine to check data validity in ARGOS message.

‘Bathy Systems, Inc. RAFOS Float data transmission.

‘3 December, 1990.

‘The 1st of 32 bytes in an ARGOS message is the CRC.

‘The function CRC will compute CRC for byte 2 through 32.

‘Hasard is used for Random because Random is reserved by BASIC.

‘Stored as file CRC in C:\RAFOS\RAF11.

DECLARE SUB Hasard (ByteN AS INTEGER)

DEFINT A-Z

DIM in(32) AS INTEGER

‘RAF11F message number 08 HEX ID 11502 01-02-93   CRC is O.K.

A$ = “8F00081C8E47239148A4D2E9743A1D0E070381C06030984C2693492492C964B2”

N = 32

FOR I = 1 to N

in(I) = VAL(“&H” + MID$(A$, 2 + I - 1, 2))

NEXT I

PRINT in(1); CRC(in(), N);

FUNCTION CRC% (IN() AS INTEGER, N AS INTEGER) STATIC

DIM ByteN as INTEGER

I = 2

ByteN = in(2)

DO

CALL Hasard(ByteN)

I = I + 1

ByteN = ByteN XOR in(I)

LOOP UNTIL  I = N

CALL Hasard (ByteN)

CRC = ByteN

END FUNCTION

DEFINT A-Z

SUB Hasard (ByteN AS INTEGER) STATIC

x% = 0

IF ByteN = 0 THEN ByteN = 127: EXIT SUB

IF (ByteN AND 1) = 1 THEN x% = x% + 1

IF (ByteN AND 4) = 4 THEN x% = x% + 1

IF (ByteN AND 8) = 8 THEN x% = x% + 1

IF (ByteN and 16) = 16 THEN x% = x% + 1

IF (X% AND 1) = 1 THEN

ByteN = INT(ByteN / 2) + 128

ELSE

ByteN = INT(ByteN / 2)

END IF

END SUB

## CONSTANTS

The pressure is measured every 6 seconds.  Temperature, salinity and pressure are measured and stored at each point in the depth table.

Two hex bytes are stored for each sensor.  The decimal numbers from the SBE sensors are converted to hex for compression in the ARGOS transmission as follows:

Temperature:	first 5 digits, 1 milli-degree resolution.

Salinity:  	5 digits starting with the second digit (first digit is usually a 3).

Pressure:  	first 5 digits, 10 cm resolution.

To convert the hex ARGOS message back to decimal numbers:

hex 	 	dec	=	converted	units

Temperature:	23CA		9162	=	9.162 		C

Salinity:	see below

Pressure:	1D4C		7500	=	750.0 		decibars

Note:  Several Salinity examples are given due to the fact that the hex range is 0000 to FFFF and the first salinity digit in the ocean is usually a 3 but can be otherwise.  Some knowledge of the area of ocean being profiled is necessary to convert the compressed data.

hex 	 	dec	=	converted

0001 		00001	=	30.0001

FFFF 		65535 	= 	36.5535 (after FFFF hex the numbers will roll over)

10001 		65537	=	36.5537 (add a 1 to the front of the rolled over hex value)

153B		05435	=	30.5435 (could be this salinity or…)

153B		05435	=	20.5435 (could be this salinity or…)

1153B		70971 =	37.0971 (could be this salinity if the numbers rolled over)

Care to inspect a graph of the converted salinities with regard to depth and temperature must be exercised in order to quality control the conversions of compressed hex to converted salinity.

Voltage (V)  =  counts/10 + .6  (counts is in decimal number) nominally 15 V and decreasing.

Vacuum (inHg)  =  counts *-0.376 + 29.15  (counts is in decimal number) nominally 5 inHg.

Depth Table 6  (in db)

Bottom, 1580, 1530, 1480, 1430, 1380, 1330, 1280, 1230, 1180, 1130, 1080, 1030, 980, 930, 880, 830, 780, 730, 680, 630, 580, 530, 480, 460, 440, 420, 400, 380, 360, 340, 320, 300, 280, 260, 240, 220, 200, 190, 180, 170, 160, 150, 140, 130, 120, 110, 100, 90, 80, 70, 60, 50, 40, 30, 20, 10, 4 or Surface.

Note:  Surface measurement has an offset of 5 db as a stop profiling point so as to leave the Sea-Bird cell full of water while transmitting.

# MISSIONS

INSTRUMENT # 52

S 2 SBE 41-ALACE V 2.0  SERIAL NO. 0120

+

H  APEX version  05 20 99  sn 0289

F3BD6 ARGOS ID number.

090  seconds repetition rate.

002  hour Trip interval.

168  intervals DOWN.

012  intervals UP.

1500  d-bar target pressure.

035  target piston position.

012  ascent rate correction factor.

024  ballast piston position.

248  piston full extention count.

068  OK vacuum count.

INSTRUMENT # 53

S 2 SBE 41-ALACE V 2.0  SERIAL NO. 0117

+

H  APEX version  05 20 99  sn 0290

F3C7E ARGOS ID number.

090  seconds repetition rate.

002  hour Trip interval.

168  intervals DOWN.

012  intervals UP.

1500  d-bar target pressure.

035  target piston position.

012  ascent rate correction factor.

024  ballast piston position.

248  piston full extention count.

068  OK vacuum count.

INSTRUMENT # 54

S 2 SBE 41-ALACE V 2.0  SERIAL NO. 0121

+

H  APEX version  05 20 99  sn 0291

F3D32 ARGOS ID number.

090  seconds repetition rate.

002  hour Trip interval.

168  intervals DOWN.

012  intervals UP.

1500  d-bar target pressure.

035  target piston position.

012  ascent rate correction factor.

024  ballast piston position.

248  piston full extention count.

067  OK vacuum count.

# RECORDS & CALIBRATIONS